## Errors by Device
Errors by Device
Date Range: -
States: Georgia, Massachusetts, Ohio, Pennsylvania
County: , Allegheny County, Cuyahoga County, Fulton County, Hampshire County, Philadelphia County, Suffolk County, Summit County, Wayne County
Device | Serial | Last Upload | Total Tests Run During Period | MTB- | MTB+ Rif- | MTB+ Rif+ | MTB+ Rif? | Invalid | Errors | Error Percentage
Wooster MicroLab | DevSerial-263 | 2018-07-09 06:09 | 2653 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 13 | 64 | 2%
Wooster Hospital | DevSerial-989 | 2018-07-10 11:27 | 44 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 0 | 0%
systemonetest | WINDOWS-7-VM | 2019-08-19 16:13 | 8 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 0 | 0%
Stefan Troubles | Tranquilo | 2018-10-03 18:07 | 63 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 1 | 1 | 2%
Shelburne Falls MicroLab | DevSerial-262 | 2018-07-09 13:31 | 3983 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 87 | 192 | 5%
Pittsburgh MicroLab | DevSerial-267 | 2018-07-06 12:56 | 24 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 1 | 4%
Philadelphia MicroLab | DevSerial-264 | 2018-07-09 11:17 | 2331 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 38 | 177 | 8%
OH Central MicroLab | DevSerial-993 | 2018-06-25 10:02 | 272 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 3 | 29 | 11%
Northampton MicroLab | DevSerial-261 | 2018-07-09 13:06 | 2949 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 28 | 184 | 6%
Cleveland MicroLab | DevSerial-265 | 2018-07-06 13:41 | 4 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 0 | 0%
Boston Micro Lab | DevSerial-260 | 2018-07-08 08:40 | 6669 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 76 | 295 | 4%
Atlanta MicroLab | DevSerial-268 | 2018-07-07 08:35 | 4 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 0 | 0%
Akron MicroLab | DevSerial-266 | 2018-07-09 09:09 | 45 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 2 | 4%
Subtotal for devices not assigned to a state |  |  | 71 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 1 | 1 | 1%
Subtotal for Georgia |  |  | 4 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 0 | 0%
Subtotal for Massachusetts |  |  | 13873 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 194 | 700 | 5%
Subtotal for Ohio |  |  | 2746 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 13 | 66 | 2%
Subtotal for Pennsylvania |  |  | 2355 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 38 | 178 | 8%
Total |  |  | 19049 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 246 | 945 | 5%

## Error Codes by Device
Error Codes by Device
Dates: -
Type of error | Wooster MicroLab | Wooster Hospital | systemonetest | Stefan Troubles | Shelburne Falls MicroLab | Pittsburgh MicroLab | Philadelphia MicroLab | OH Central MicroLab | Northampton MicroLab | Cleveland MicroLab | Boston Micro Lab | Atlanta MicroLab | Akron MicroLab
5011 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5007 | 271 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5006 | 66 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2008 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2037 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2025 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2022 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Other Error | 40 | 10 | 7 | 60 | 373 | 294 | 16 | 66 | 36 | 6 | 126 | 120 | 144
Invalid Results | 155 | 12 | 0 | 64 | 29 | 73 | 5 | 73 | 12 | 3 | 13 | 15 | 11
Total errors, invalid results, no results | 570 | 24 | 7 | 124 | 402 | 367 | 21 | 139 | 48 | 9 | 139 | 135 | 155

## Error Codes
Code | Category | Description
1001 | Run-time error | The actual temperature has drifted too far away from the setpoint
1002 | Run-time error | The temperature difference exceeds the limit
1004 | Run-time error | The internal instrument temperature was out of range
1005 | Run-time error | Optic signal of from detector exceeded the limit
1006 | Run-time error | Detector dark signal exceeded the limit
1007 | Run-time error | The power supply was detected to be out of range
1017 | Run-time error | The measured temperature of the optical system was not within the acceptable range
1018 | Run-time error | A valve positioning error was detected at the end of the run
2003 | Operation terminated error | Module is already running a test while performing command
2005 | Operation terminated error | Motion of the syringe drive was not detected
2008 | Operation terminated error | Syringe pressure reading exceeds the protocol limit
2009 | Operation terminated error | Syringe pressure reading is below the protocol limit
2011 | Cartridge loading error | Unable to initialize pressure sensor
2014 | Operation terminated error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
2016 | Operation terminated error | The system was unable to find the valve home position
2018 | Cartridge loading error | Attempt to load a cartridge while the door is still closed
2022 | Operation terminated error | Failed to get to desired temperature
2025 | Cartridge loading error | The system failed to find the plunger home position
2032 | Operation terminated error | The ultrasonic horn could not be tuned properly
2034 | Operation terminated error | The optical signal from Detector did not reach the expected value
2037 | Cartridge loading error | The cartridge integrity test failed
4010 | Self-test error | Cooling fan operation was not verified. Measured temperature exceeded the limit
4017 | Self-test error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
5006 | Post-run analysis error | Probe check failed. Probe check value was above the maximum
5007 | Post-run analysis error | Probe check failed. Probe check value was below the minimum
5011 | Post-run analysis error | Signal loss detected in the amplification curve