## Errors by Device
<Title>
<DateRange>
<States>
<County>
<Table>
<Device> | <Serial> | <LastUpload> | <TotalTests> | MTB- | MTB+ Rif- | MTB+ Rif+ | MTB+ Rif? | <Invalid> | <Errors> | <ErrorPercentage>
<deviceName> | <serial> | <lastUpload> | <totalTests> | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | <invalid> | <errors> | <errorPercentage>
<SubtotalNotAssigned> |  |  | <#> | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | <#> | <#> | #VALUE!
<SubtotalAssigned> |  |  | <#> | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | <#> | <#> | #VALUE!
<Total> |  |  | <#> | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | <#> | <#> | #VALUE!

## Error Codes by Device
Error Codes by Device
Dates: -
Type of error | Wooster MicroLab | Wooster Hospital | systemonetest | Stefan Troubles | Shelburne Falls MicroLab | Pittsburgh MicroLab | Philadelphia MicroLab | OH Central MicroLab | Northampton MicroLab | Cleveland MicroLab | Boston Micro Lab | Atlanta MicroLab | Akron MicroLab
5011 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5007 | 271 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5006 | 66 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2008 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2037 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2025 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2022 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Other Error | 40 | 10 | 7 | 60 | 373 | 294 | 16 | 66 | 36 | 6 | 126 | 120 | 144
Invalid Results | 155 | 12 | 0 | 64 | 29 | 73 | 5 | 73 | 12 | 3 | 13 | 15 | 11
Total errors, invalid results, no results | 570 | 24 | 7 | 124 | 402 | 367 | 21 | 139 | 48 | 9 | 139 | 135 | 155

## Error Codes
Code | Category | Description
1001 | Run-time error | The actual temperature has drifted too far away from the setpoint
1002 | Run-time error | The temperature difference exceeds the limit
1004 | Run-time error | The internal instrument temperature was out of range
1005 | Run-time error | Optic signal of from detector exceeded the limit
1006 | Run-time error | Detector dark signal exceeded the limit
1007 | Run-time error | The power supply was detected to be out of range
1017 | Run-time error | The measured temperature of the optical system was not within the acceptable range
1018 | Run-time error | A valve positioning error was detected at the end of the run
2003 | Operation terminated error | Module is already running a test while performing command
2005 | Operation terminated error | Motion of the syringe drive was not detected
2008 | Operation terminated error | Syringe pressure reading exceeds the protocol limit
2009 | Operation terminated error | Syringe pressure reading is below the protocol limit
2011 | Cartridge loading error | Unable to initialize pressure sensor
2014 | Operation terminated error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
2016 | Operation terminated error | The system was unable to find the valve home position
2018 | Cartridge loading error | Attempt to load a cartridge while the door is still closed
2022 | Operation terminated error | Failed to get to desired temperature
2025 | Cartridge loading error | The system failed to find the plunger home position
2032 | Operation terminated error | The ultrasonic horn could not be tuned properly
2034 | Operation terminated error | The optical signal from Detector did not reach the expected value
2037 | Cartridge loading error | The cartridge integrity test failed
4010 | Self-test error | Cooling fan operation was not verified. Measured temperature exceeded the limit
4017 | Self-test error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
5006 | Post-run analysis error | Probe check failed. Probe check value was above the maximum
5007 | Post-run analysis error | Probe check failed. Probe check value was below the minimum
5011 | Post-run analysis error | Signal loss detected in the amplification curve