## Errors by Module
Erreurs par module
Dates: <Dates>
<StatesLabel>: <States>
<LgasLabel>: <Lgas>
<SitesLabel>: <Sites>
<Table>
# | Nom du déploiement | Zone | District | InstrumentSérie | ModuleSérie | Nom du module | Date de l'erreur | Code d'erreur | Nombre total de tests exécutés pendant la période | Nombre total d'erreurs pendant la période
<#> | <deploymentName> | <zone> | <district> | <instrumentSerial> | <moduleSerial> | <moduleName> | <dateOfError> | <errorCode> | <totalTests> | <totalErrors>

## Error Codes
Code | Category | Description
1001 | Run-time error | The actual temperature has drifted too far away from the setpoint
1002 | Run-time error | The temperature difference exceeds the limit
1004 | Run-time error | The internal instrument temperature was out of range
1005 | Run-time error | Optic signal of from detector exceeded the limit
1006 | Run-time error | Detector dark signal exceeded the limit
1007 | Run-time error | The power supply was detected to be out of range
1017 | Run-time error | The measured temperature of the optical system was not within the acceptable range
1018 | Run-time error | A valve positioning error was detected at the end of the run
2003 | Operation terminated error | Module is already running a test while performing command
2005 | Operation terminated error | Motion of the syringe drive was not detected
2008 | Operation terminated error | Syringe pressure reading exceeds the protocol limit
2009 | Operation terminated error | Syringe pressure reading is below the protocol limit
2011 | Cartridge loading error | Unable to initialize pressure sensor
2014 | Operation terminated error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
2016 | Operation terminated error | The system was unable to find the valve home position
2018 | Cartridge loading error | Attempt to load a cartridge while the door is still closed
2022 | Operation terminated error | Failed to get to desired temperature
2025 | Cartridge loading error | The system failed to find the plunger home position
2032 | Operation terminated error | The ultrasonic horn could not be tuned properly
2034 | Operation terminated error | The optical signal from Detector did not reach the expected value
2037 | Cartridge loading error | The cartridge integrity test failed
4010 | Self-test error | Cooling fan operation was not verified. Measured temperature exceeded the limit
4017 | Self-test error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
5006 | Post-run analysis error | Probe check failed. Probe check value was above the maximum
5007 | Post-run analysis error | Probe check failed. Probe check value was below the minimum
5011 | Post-run analysis error | Signal loss detected in the amplification curve