## Errors by Device
Errors by Device
Dates: <Dates>
<StatesLabel>: <States>
<LgasLabel>: <Lgas>
<SitesLabel>: <Sites>
<Table> | [6 empty] | Same origin site in Test and Sample
Device | Serial | Last Upload | Total Tests Run During Period | # Invalid | # Errors | Error Percentage | Total Tests Run During Period | # Invalid | # Errors | Error Percentage
<name> | <serial> | <lastUploadOn> | <totalResults> | <totalInvalid> | <totalErrors> | <errorRate> | <totalResults_sameOriginSite> | <totalInvalid_sameOriginSite> | <totalErrors_sameOriginSite> | <errorRate_sameOriginSite>

## Error Codes by Device
Error Codes by Device
Dates: <Dates>
<StatesLabel>: <States>
<LgasLabel>: <Lgas>
<SitesLabel>: <Sites>
<Table>
Type of error | 5011 | 5007 | 5006 | 2008 | 2005 | 2037 | 2025 | 2022 | Other Error | Invalid Results | No Results | Total errors, invalid results, no results
<deviceName> | <5011> | <5007> | <5006> | <2008> | <2005> | <2037> | <2025> | <2022> | <otherErrors> | <invalidResults> | <noResults> | #VALUE!

## Error Codes
Code | Category | Description
1001 | Run-time error | The actual temperature has drifted too far away from the setpoint
1002 | Run-time error | The temperature difference exceeds the limit
1004 | Run-time error | The internal instrument temperature was out of range
1005 | Run-time error | Optic signal of from detector exceeded the limit
1006 | Run-time error | Detector dark signal exceeded the limit
1007 | Run-time error | The power supply was detected to be out of range
1017 | Run-time error | The measured temperature of the optical system was not within the acceptable range
1018 | Run-time error | A valve positioning error was detected at the end of the run
2003 | Operation terminated error | Module is already running a test while performing command
2005 | Operation terminated error | Motion of the syringe drive was not detected
2008 | Operation terminated error | Syringe pressure reading exceeds the protocol limit
2009 | Operation terminated error | Syringe pressure reading is below the protocol limit
2011 | Cartridge loading error | Unable to initialize pressure sensor
2014 | Operation terminated error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
2016 | Operation terminated error | The system was unable to find the valve home position
2018 | Cartridge loading error | Attempt to load a cartridge while the door is still closed
2022 | Operation terminated error | Failed to get to desired temperature
2025 | Cartridge loading error | The system failed to find the plunger home position
2032 | Operation terminated error | The ultrasonic horn could not be tuned properly
2034 | Operation terminated error | The optical signal from Detector did not reach the expected value
2037 | Cartridge loading error | The cartridge integrity test failed
4010 | Self-test error | Cooling fan operation was not verified. Measured temperature exceeded the limit
4017 | Self-test error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
5006 | Post-run analysis error | Probe check failed. Probe check value was above the maximum
5007 | Post-run analysis error | Probe check failed. Probe check value was below the minimum
5011 | Post-run analysis error | Signal loss detected in the amplification curve