## Devices
GeneXpert Devices
Dates: <Dates> |  | Notes
<StatesLabel>: <States> |  | Please note that this report includes MTB/Rif, MTB ULTRA and MTB XDR results
<LgasLabel>: <Lgas>
<SitesLabel>: <Sites>
<Table>
Device | State | District | Serial | Instrument Serial | Last Upload | Status | Status Note | Last Ping On | Last Ping From | Information Below as of | Operating System | Time Zone | Language / Region | GxConnect Version | Device Found? | LIS Configured Correctly? | GeneXpert Installation Date | Total | Other
<name> | <stateName> | <districtName> | <serial> | <instrumentSerial> | <lastUploadOn> | <status> | <statusNote> | <lastPingOn> | <lastPingFrom> | <informationDate> | <operativeSystem> | <timeZone> | <language> | <gxConnectVersion> | <deviceFound> | <lisConfiguredCorrectly> | <instalationDate> | <totalXpertTests> | <totalOtherTests>

## Inventory
Inventory
Date Range: -
States: , Berea, Butha-Buthe, Leribe, Mafeteng, Maseru, Mohale's Hoek, Mokhotlong, Qacha's Nek, Quthing, Thaba-Tseka
Deployment | Serial | Last Upload | Total Tests | Average Consumption rate | Utilization Rate* | Projected annual cartridge consumption | Total Inventory | Days of Inventory
BEREA LAB TB | CPHD-0F06604 | 2022-07-25 12:28 | 9609 |  |  | 0 | 0 | 0
BUTHA-BUTHE HOSPITAL LAB TB | Cepheid-3200YJJ | 2022-07-25 11:17 | 5518 |  |  | 0 | 0 | 0
Katse mini lab TB | Cepheid-6131P1X | 2022-07-15 11:52 | 114 |  |  | 0 | 0 | 0
Machabeng Hospital | Cepheid-3381P57 | 2022-07-22 13:39 | 2165 |  |  | 0 | 0 | 0
MAFETENG HOSPITAL LAB | CPHD-0F13818 | 2022-07-15 12:13 | 4579 |  |  | 0 | 0 | 0
MAFETENG HOSPITAL MCH | Cepheid-62952KF | 2021-08-13 10:49 | 4 |  |  | 0 | 0 | 0
MAKOANYANE HOSPITAL LAB | Cepheid-70310VQ | 2022-07-25 10:33 | 1656 |  |  | 0 | 0 | 0
MALUTI LABORATORY TB | Cepheid-2UA3481GQ5 | 2022-07-25 10:40 | 9103 |  |  | 0 | 0 | 0
MAMOHAU HOSPITAL LABORATORY TB | Cepheid-44224V8 | 2022-07-25 11:00 | 7193 |  |  | 0 | 0 | 0
MAPUTSOE FILTER CLINIC | Cepheid-6131P1D | 2022-07-08 12:21 | 2530 |  |  | 0 | 0 | 0
MOHALE MINI LAB | Cepheid-6131P2G | 2022-07-21 11:05 | 11 |  |  | 0 | 0 | 0
MOKHOTLONG LABORATORY TB | G2G0LS1 | 2022-07-22 18:27 | 4310 |  |  | 0 | 0 | 0
Mokhotlong Mapholaneng mini lab TB | Cepheid-6131P2D | 2022-07-22 08:50 | 786 |  |  | 0 | 0 | 0
MOTEBANG HOSPITAL LAB | Cepheid-44224VF | 2022-07-21 09:16 | 14828 |  |  | 0 | 0 | 0
NTRL-16 | Cepheid-7361P5F | 2020-09-16 12:29 | 4789 |  |  | 0 | 0 | 0
NTRL-4 | D1L9ZKN1 | 2022-07-22 12:07 | 1449 |  |  | 0 | 0 | 0
PARAY HOSPITAL LABORATORY TB | CPHD-0F13805 | 2022-03-02 14:57 | 3883 |  |  | 0 | 0 | 0
PIH BOBETE | Cepheid-6201QKC | 2022-07-25 10:26 | 1151 |  |  | 0 | 0 | 0
PIH LAB BOTSHABELO TB | 5CG5010C6L | 2022-07-25 11:58 | 15192 |  |  | 0 | 0 | 0
PIH LEBAKENG | Cepheid-6131P2X | 2022-05-24 12:25 | 799 |  |  | 0 | 0 | 0
PIH LESOTHO | CPHD-0F10856 | 2022-04-14 05:43 | 54 |  |  | 0 | 0 | 0
PIH NKAU | Cepheid-6131P3F | 2022-07-22 15:58 | 730 |  |  | 0 | 0 | 0
PIH NOHANA | Cepheid-6201QKK | 2022-05-04 18:00 | 949 |  |  | 0 | 0 | 0
QUEEN II 4 MODULES | Cepheid-7282LQK | 2022-07-20 12:31 | 16178 |  |  | 0 | 0 | 0
Quthing Lab | Cepheid-44224T0 | 2022-07-25 11:31 | 12116 |  |  | 0 | 0 | 0
Scott Lab TB | CPHD-0F08704 | 2022-07-22 12:37 | 5869 |  |  | 0 | 0 | 0
Sehlabathebe Mini Lab | Cepheid-6131P3Z | 2022-04-06 10:07 | 38 |  |  | 0 | 0 | 0
Sehong-Hong Thaba-Tseka | Cepheid-6131P32 | 2022-02-08 06:39 | 119 |  |  | 0 | 0 | 0
Semonkong mini lab TB | Cepheid-6021ZML | 2022-07-25 11:07 | 768 |  |  | 0 | 0 | 0
ST JAMES | Cepheid-44224T7 | 2022-07-22 12:10 | 345 |  |  | 0 | 0 | 0
St Joseph's Roma | CPHD-0F09918 | 2022-07-25 06:49 | 1311 |  |  | 0 | 0 | 0
St Matthews H/C Quthing | Cepheid-3481GRC | 2022-07-13 09:18 | 2347 |  |  | 0 | 0 | 0
TEBA LERIBE LAB TB | Cepheid-3031JBL | 2021-06-07 12:40 | 2795 |  |  | 0 | 0 | 0
Teba Mafeteng Lab TB | Cepheid2FSQ0R1 | 2021-06-21 13:26 | 2216 |  |  | 0 | 0 | 0
Tebellong Hospital | Cepheid-33714J3 | 2022-07-22 11:57 | 796 |  |  | 0 | 0 | 0
THABANA-MORENA HEALTH CENTRE TB | CPHD-0F07064 | 2022-07-21 14:13 | 318 |  |  | 0 | 0 | 0
Subtotal for devices not assigned to a state |  |  | 12518 | 0.00 |  | 0 | 0
Subtotal for Berea |  |  | 18712 | 0.00 |  | 0 | 0
Subtotal for Butha-Buthe |  |  | 5518 | 0.00 |  | 0 | 0
Subtotal for Leribe |  |  | 14828 | 0.00 |  | 0 | 0
Subtotal for Mafeteng |  |  | 7117 | 0.00 |  | 0 | 0
Subtotal for Maseru |  |  | 47277 | 0.00 |  | 0 | 0
Subtotal for Mohale's Hoek |  |  | 730 | 0.00 |  | 0 | 0
Subtotal for Mokhotlong |  |  | 5096 | 0.00 |  | 0 | 0
Subtotal for Qacha's Nek |  |  | 4747 | 0.00 |  | 0 | 0
Subtotal for Quthing |  |  | 14463 | 0.00 |  | 0 | 0
Subtotal for Thaba-Tseka |  |  | 5612 | 0.00 |  | 0 | 0
Total |  |  | 136618 | 0.00 |  | 0 | 0
*assumes 4-module GeneXpert running 8 hours a day, 2 hours per test or 16 tests/day
**calculated based on average consumption rate

## Devices TB Reach
List of Devices for TB Reach
Dates: -
States: , Berea, Butha-Buthe, Leribe, Mafeteng, Maseru, Mohale's Hoek, Mokhotlong, Qacha's Nek, Quthing, Thaba-Tseka
| BEREA LAB TB | BUTHA-BUTHE HOSPITAL LAB TB | Katse mini lab TB | Machabeng Hospital | MAFETENG HOSPITAL LAB | MAFETENG HOSPITAL MCH | MAKOANYANE HOSPITAL LAB | MALUTI LABORATORY TB | MAMOHAU HOSPITAL LABORATORY TB | MAPUTSOE FILTER CLINIC | MOHALE MINI LAB | MOKHOTLONG LABORATORY TB | Mokhotlong Mapholaneng mini lab TB | MOTEBANG HOSPITAL LAB | NTRL-16 | NTRL-4 | PARAY HOSPITAL LABORATORY TB | PIH BOBETE | PIH LAB BOTSHABELO TB | PIH LEBAKENG | PIH LESOTHO | PIH NKAU | PIH NOHANA | QUEEN II 4 MODULES | Quthing Lab | Scott Lab TB | Sehlabathebe Mini Lab | Sehong-Hong Thaba-Tseka | Semonkong mini lab TB | ST JAMES | St Joseph's Roma | St Matthews H/C Quthing | TEBA LERIBE LAB TB | Teba Mafeteng Lab TB | Tebellong Hospital | THABANA-MORENA HEALTH CENTRE TB
BK total
BK (cn, rec)
BK+
BK-
%BK+ | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
Xpert Tests | 9335 | 5330 | 109 | 2103 | 4391 | 4 | 1601 | 8895 | 6934 | 2499 | 11 | 4182 | 762 | 14327 | 4706 | 1411 | 3798 | 1127 | 14911 | 744 | 49 | 707 | 935 | 15621 | 11714 | 5698 | 38 | 114 | 745 | 329 | 1262 | 2301 | 2736 | 2133 | 758 | 306
%Xpert Tests | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
Xpert+ | 808 | 421 | 10 | 114 | 236 | 0 | 88 | 594 | 177 | 117 | 0 | 134 | 34 | 772 | 155 | 69 | 241 | 24 | 438 | 15 | 14 | 52 | 23 | 1342 | 866 | 387 | 2 | 10 | 39 | 12 | 94 | 112 | 59 | 67 | 78 | 35
Xpert- | 8272 | 4711 | 93 | 1905 | 4039 | 4 | 1424 | 8163 | 6473 | 2326 | 11 | 3945 | 704 | 13048 | 4469 | 1266 | 3449 | 1066 | 13750 | 671 | 27 | 620 | 891 | 13772 | 10365 | 5121 | 36 | 90 | 639 | 303 | 1069 | 2121 | 2595 | 1988 | 623 | 256
Xpert Error/Invalid | 209 | 168 | 0 | 73 | 105 | 0 | 74 | 92 | 260 | 43 | 0 | 89 | 19 | 456 | 51 | 43 | 95 | 31 | 458 | 53 | 2 | 27 | 14 | 402 | 420 | 176 | 0 | 9 | 59 | 11 | 84 | 50 | 77 | 69 | 40 | 10
%Xpert+ | 8.656% | 7.899% | 9.174% | 5.421% | 5.375% | 0% | 5.497% | 6.678% | 2.553% | 4.682% | 0% | 3.204% | 4.462% | 5.388% | 3.294% | 4.89% | 6.345% | 2.13% | 2.937% | 2.016% | 28.57% | 7.355% | 2.46% | 8.591% | 7.393% | 6.792% | 5.263% | 8.772% | 5.235% | 3.647% | 7.448% | 4.867% | 2.156% | 3.141% | 10.29% | 11.44%
%Xpert Increase | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
%Xpert Error/Invalid | 2.239% | 3.152% | 0% | 3.471% | 2.391% | 0% | 4.622% | 1.034% | 3.75% | 1.721% | 0% | 2.128% | 2.493% | 3.183% | 1.084% | 3.047% | 2.501% | 2.751% | 3.072% | 7.124% | 4.082% | 3.819% | 1.497% | 2.573% | 3.585% | 3.089% | 0% | 7.895% | 7.919% | 3.343% | 6.656% | 2.173% | 2.814% | 3.235% | 5.277% | 3.268%
Xpert RIF+ | 38 | 29 | 6 | 8 | 10 | 0 | 15 | 40 | 23 | 10 | 0 | 13 | 5 | 44 | 26 | 33 | 13 | 6 | 258 | 5 | 5 | 7 | 6 | 86 | 59 | 12 | 0 | 5 | 8 | 3 | 15 | 14 | 4 | 7 | 15 | 5
Xpert RIF IND | 8 | 1 | 0 | 3 | 1 | 0 | 0 | 6 | 1 | 3 | 0 | 1 | 0 | 7 | 5 | 0 | 0 | 0 | 7 | 0 | 1 | 1 | 1 | 19 | 4 | 2 | 0 | 0 | 0 | 0 | 0 | 4 | 1 | 2 | 2 | 0
Xpert+ Treated
%Xpert+ Treated | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
Xpert per day | 424.3182 | 242.2727 | 4.9545 | 95.5909 | 199.5909 | 0.18182 | 72.7727 | 404.3182 | 315.1818 | 113.5909 | 0.5 | 190.0909 | 34.6364 | 651.2273 | 213.9091 | 64.1364 | 172.6364 | 51.2273 | 677.7727 | 33.8182 | 2.2273 | 32.1364 | 42.5 | 710.0455 | 532.4545 | 259 | 1.7273 | 5.1818 | 33.8636 | 14.9545 | 57.3636 | 104.5909 | 124.3636 | 96.9545 | 34.4545 | 13.9091
Type of errors
5011 | 9 | 9 | 0 | 19 | 7 | 0 | 6 | 8 | 11 | 5 | 0 | 1 | 0 | 15 | 6 | 0 | 1 | 1 | 41 | 2 | 0 | 1 | 0 | 13 | 24 | 7 | 0 | 1 | 0 | 0 | 2 | 3 | 1 | 3 | 8 | 0
5007 | 64 | 63 | 0 | 24 | 20 | 0 | 7 | 30 | 56 | 9 | 0 | 47 | 7 | 164 | 9 | 3 | 25 | 5 | 149 | 5 | 0 | 4 | 2 | 158 | 77 | 30 | 0 | 1 | 4 | 2 | 12 | 11 | 24 | 36 | 11 | 1
5006 | 4 | 12 | 0 | 1 | 4 | 0 | 1 | 8 | 22 | 3 | 0 | 4 | 1 | 87 | 2 | 0 | 2 | 1 | 35 | 1 | 0 | 1 | 0 | 58 | 15 | 13 | 0 | 0 | 3 | 0 | 38 | 3 | 1 | 3 | 1 | 0
2008 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2037 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2025 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2022 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Other Error | 112 | 35 | 0 | 22 | 70 | 0 | 19 | 27 | 123 | 14 | 0 | 17 | 6 | 113 | 14 | 26 | 50 | 17 | 174 | 41 | 2 | 19 | 10 | 126 | 245 | 103 | 0 | 6 | 50 | 7 | 26 | 23 | 31 | 19 | 12 | 9
Invalid Results | 20 | 49 | 0 | 7 | 4 | 0 | 41 | 19 | 48 | 12 | 0 | 20 | 5 | 77 | 20 | 14 | 17 | 7 | 59 | 4 | 0 | 2 | 2 | 47 | 59 | 23 | 0 | 1 | 2 | 2 | 6 | 10 | 20 | 8 | 8 | 0
Total errors, invalid results, no results | 209 | 168 | 0 | 73 | 105 | 0 | 74 | 92 | 260 | 43 | 0 | 89 | 19 | 456 | 51 | 43 | 95 | 31 | 458 | 53 | 2 | 27 | 14 | 402 | 420 | 176 | 0 | 9 | 59 | 11 | 84 | 50 | 77 | 69 | 40 | 10

## Error Codes
Code | Category | Description
1001 | Run-time error | The actual temperature has drifted too far away from the setpoint
1002 | Run-time error | The temperature difference exceeds the limit
1004 | Run-time error | The internal instrument temperature was out of range
1005 | Run-time error | Optic signal of from detector exceeded the limit
1006 | Run-time error | Detector dark signal exceeded the limit
1007 | Run-time error | The power supply was detected to be out of range
1017 | Run-time error | The measured temperature of the optical system was not within the acceptable range
1018 | Run-time error | A valve positioning error was detected at the end of the run
2003 | Operation terminated error | Module is already running a test while performing command
2005 | Operation terminated error | Motion of the syringe drive was not detected
2008 | Operation terminated error | Syringe pressure reading exceeds the protocol limit
2009 | Operation terminated error | Syringe pressure reading is below the protocol limit
2011 | Cartridge loading error | Unable to initialize pressure sensor
2014 | Operation terminated error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
2016 | Operation terminated error | The system was unable to find the valve home position
2018 | Cartridge loading error | Attempt to load a cartridge while the door is still closed
2022 | Operation terminated error | Failed to get to desired temperature
2025 | Cartridge loading error | The system failed to find the plunger home position
2032 | Operation terminated error | The ultrasonic horn could not be tuned properly
2034 | Operation terminated error | The optical signal from Detector did not reach the expected value
2037 | Cartridge loading error | The cartridge integrity test failed
4010 | Self-test error | Cooling fan operation was not verified. Measured temperature exceeded the limit
4017 | Self-test error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
5006 | Post-run analysis error | Probe check failed. Probe check value was above the maximum
5007 | Post-run analysis error | Probe check failed. Probe check value was below the minimum
5011 | Post-run analysis error | Signal loss detected in the amplification curve