## Errors by Device
Failed Test Report
Deployment | Last Upload | Total Tests Run During Period | MTB- | MTB+ Rif- | MTB+ Rif+ | MTB+ Rif? | Invalid | Errors | Error Percentage
Wooster CTRL | 2018-02-15 11:27 | 44 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 0 | 0%
Wooster Community | 2018-02-14 06:09 | 2653 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 13 | 64 | 2%
Shelburne Falls Lab | 2018-02-15 13:31 | 3983 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 87 | 192 | 5%
Pittsburgh Lab | 2018-02-07 12:56 | 24 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 0 | 1 | 4%
Philadelphia Lab I | 2018-02-15 11:17 | 2331 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 38 | 177 | 8%
Northampton Hospital | 2018-02-15 13:06 | 2950 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 28 | 184 | 6%
Subtotal for Massachusetts |  | 13600 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 191 | 671 | 5%
Subtotal for Ohio |  | 2702 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 13 | 66 | 2%
Subtotal for Pennsylvania |  | 2355 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 38 | 178 | 8%
Total |  | 29974 | <MtbNegative> | <MtbPositiveRifNegative> | <MtbPositiveRifPositive> | <MtbPositiveRifIndeterminate> | 343 | 1252 | 4%

## Error Codes by Device
Error Codes by Device
Dates: -
States: Georgia, Massachusetts, Ohio, Pennsylvania
County: , Allegheny County, Cuyahoga County, Fulton County, Hampshire County, Philadelphia County, Suffolk County, Summit County, Wayne County
Type of error | Zankli Medical Cente | Worawora | Wooster CTRL | Wooster Community | VRH,Ho | systemonetest | SSH Ikole Ekiti | Sogakope | Shelburne Falls Lab | Pittsburgh Lab | Philadelphia Lab I | Papase | Northampton Hospital | Nkwanta, VR | Nkawie-Toase | LIS | LIS | Kumasi South | Kuje GH | Kpassa | Kpando | Keta | KCCR | KATH | ISSH Owerri | INFECTIOUS DISEASE H | Hohoe | GH Otukpo | GH KUJE | GH KATSINA-ALA | GH Kagarko | GH IKONO | Genexpert | General Hospital Kon | GENERAL HOSPITAL KAU | General Hospital Iqu | FMC Makurdi | Dzodze | CTRL LAB | Cleveland Lab | CHC AWA | Boston Micro Labs | Atlanta Lab | Ankaase Meth Hosp | Akron Labs | Aflao | 68 NARH Yaba
5011 | 1 | 2 | 0 | 2 | 3 | 0 | 0 | 1 | 11 | 0 | 25 | 1 | 3 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 75 | 0 | 0 | 1 | 0 | 1
5007 | 3 | 1 | 0 | 45 | 5 | 0 | 10 | 2 | 98 | 0 | 65 | 0 | 127 | 0 | 0 | 115 | 2 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 10 | 1 | 2 | 0 | 0 | 115 | 0 | 1 | 1 | 4 | 2
5006 | 0 | 0 | 0 | 11 | 4 | 0 | 0 | 0 | 16 | 1 | 20 | 0 | 15 | 0 | 1 | 3 | 0 | 1 | 0 | 0 | 5 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 0 | 0 | 0 | 1 | 0
2008 | 0 | 13 | 0 | 6 | 24 | 0 | 0 | 4 | 49 | 0 | 28 | 2 | 33 | 1 | 0 | 1 | 3 | 0 | 0 | 10 | 2 | 9 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 1 | 0 | 0 | 0 | 78 | 0 | 0 | 0 | 0 | 0
2005 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 11 | 1 | 4 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0
2037 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2025 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2022 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 0 | 13 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Other Error | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 3 | 5 | 0 | 15 | 0 | 2 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 4 | 2 | 17 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 3 | 8 | 0 | 2 | 0 | 1 | 0
Invalid Results | 1 | 6 | 0 | 13 | 21 | 0 | 1 | 1 | 87 | 0 | 38 | 1 | 28 | 5 | 1 | 36 | 8 | 0 | 0 | 0 | 0 | 1 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 4 | 6 | 0 | 0 | 1 | 76 | 0 | 0 | 0 | 1 | 0
Total errors, invalid results, no results | 5 | 22 | 0 | 77 | 60 | 0 | 11 | 12 | 279 | 1 | 215 | 5 | 212 | 6 | 2 | 162 | 15 | 1 | 0 | 11 | 11 | 14 | 21 | 1 | 0 | 6 | 4 | 0 | 2 | 4 | 1 | 0 | 8 | 0 | 5 | 2 | 15 | 13 | 2 | 0 | 4 | 371 | 0 | 3 | 2 | 7 | 3

## Error Codes
Code | Category | Description
1001 | Run-time error | The actual temperature has drifted too far away from the setpoint
1002 | Run-time error | The temperature difference exceeds the limit
1004 | Run-time error | The internal instrument temperature was out of range
1005 | Run-time error | Optic signal of from detector exceeded the limit
1006 | Run-time error | Detector dark signal exceeded the limit
1007 | Run-time error | The power supply was detected to be out of range
1017 | Run-time error | The measured temperature of the optical system was not within the acceptable range
1018 | Run-time error | A valve positioning error was detected at the end of the run
2003 | Operation terminated error | Module is already running a test while performing command
2005 | Operation terminated error | Motion of the syringe drive was not detected
2008 | Operation terminated error | Syringe pressure reading exceeds the protocol limit
2009 | Operation terminated error | Syringe pressure reading is below the protocol limit
2011 | Cartridge loading error | Unable to initialize pressure sensor
2014 | Operation terminated error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
2016 | Operation terminated error | The system was unable to find the valve home position
2018 | Cartridge loading error | Attempt to load a cartridge while the door is still closed
2022 | Operation terminated error | Failed to get to desired temperature
2025 | Cartridge loading error | The system failed to find the plunger home position
2032 | Operation terminated error | The ultrasonic horn could not be tuned properly
2034 | Operation terminated error | The optical signal from Detector did not reach the expected value
2037 | Cartridge loading error | The cartridge integrity test failed
4010 | Self-test error | Cooling fan operation was not verified. Measured temperature exceeded the limit
4017 | Self-test error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
5006 | Post-run analysis error | Probe check failed. Probe check value was above the maximum
5007 | Post-run analysis error | Probe check failed. Probe check value was below the minimum
5011 | Post-run analysis error | Signal loss detected in the amplification curve