## Xpert Site Summary Report
Xpert Site Summary Report
Dates: <Dates> |  |  |  | Note: The result categories are represented by all the possible result combinations that GeneXpert sends to Aspect. For example, MTB Detected does not include MTB Detected Rif Detected, as they are considered different results
<StatesLabel>: <States>
<LgasLabel>: <Lgas>
<SitesLabel>: <Sites>
<AssaysLabel>: <Assays>
<Table>
Device | State | District | Serial | Instrument Serial | Last Result Update | Status | Status Notes | Last Ping On | Last Ping From | Information Below as of | Operating System | Time Zone | Language / Region | GxConnect Version | Device Found? | LIS Configured Correctly? | GeneXpert Installation Date | Total | Other
<name> | <stateName> | <districtName> | <serial> | <instrumentSerial> | <lastUploadOn> | <status> | <statusNote> | <lastPingOn> | <lastPingFrom> | <informationDate> | <operativeSystem> | <timeZone> | <language> | <gxConnectVersion> | <deviceFound> | <lisConfiguredCorrectly> | <instalationDate> | <totalXpertTests> | <totalOtherTests>

## Xpert Site Inventory
Xpert Site Inventory
Dates: <Dates> |  | Note: this report only includes the consumables that were sometime tracked on the sites
<StatesLabel>: <States>
<LgasLabel>: <Lgas>
<SitesLabel>: <Sites>
<AssaysLabel>: <Assays>
<Table>
Site | Consumable | Daily Consumption Rate | Projected Annual Consumption | Quantity On Hand | Quantity On Hand but Expired | Days Stock Left | Projected Waste
<siteName> | <consumableName> | <dailyConsumptionRate> | <projectedAnnualConsumption> | <quantityOnHand> | <quantityExpired> | <daysStockLeft> | <projectedWaste>

## Xpert Site Errors by Device
Xpert Site Errors by Devices
Dates: <Dates>
<StatesLabel>: <States>
<LgasLabel>: <Lgas>
<SitesLabel>: <Sites>
<AssaysLabel>: <Assays>
<Table> | 2
Device | Type of error | [8 empty] | Invalid results, no results | Total errors, invalid results, no results
| 5011 | 5007 | 5006 | 2008 | 2005 | 2037 | 2025 | 2022 | Other Error
<deviceName> | <5011> | <5007> | <5006> | <2008> | <2005> | <2037> | <2025> | <2022> | <otherErrors> | <invalidResults> | #VALUE!

## Error Codes
Code | Category | Description
1001 | Run-time error | The actual temperature has drifted too far away from the setpoint
1002 | Run-time error | The temperature difference exceeds the limit
1004 | Run-time error | The internal instrument temperature was out of range
1005 | Run-time error | Optic signal of from detector exceeded the limit
1006 | Run-time error | Detector dark signal exceeded the limit
1007 | Run-time error | The power supply was detected to be out of range
1017 | Run-time error | The measured temperature of the optical system was not within the acceptable range
1018 | Run-time error | A valve positioning error was detected at the end of the run
2003 | Operation terminated error | Module is already running a test while performing command
2005 | Operation terminated error | Motion of the syringe drive was not detected
2008 | Operation terminated error | Syringe pressure reading exceeds the protocol limit
2009 | Operation terminated error | Syringe pressure reading is below the protocol limit
2011 | Cartridge loading error | Unable to initialize pressure sensor
2014 | Operation terminated error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
2016 | Operation terminated error | The system was unable to find the valve home position
2018 | Cartridge loading error | Attempt to load a cartridge while the door is still closed
2022 | Operation terminated error | Failed to get to desired temperature
2025 | Cartridge loading error | The system failed to find the plunger home position
2032 | Operation terminated error | The ultrasonic horn could not be tuned properly
2034 | Operation terminated error | The optical signal from Detector did not reach the expected value
2037 | Cartridge loading error | The cartridge integrity test failed
4010 | Self-test error | Cooling fan operation was not verified. Measured temperature exceeded the limit
4017 | Self-test error | The digital temperature reading for Thermistor A/Thermistor B/Ambient Thermistor/Optic Thermistor was not within the acceptable range
5006 | Post-run analysis error | Probe check failed. Probe check value was above the maximum
5007 | Post-run analysis error | Probe check failed. Probe check value was below the minimum
5011 | Post-run analysis error | Signal loss detected in the amplification curve